## Sheet1
AuralAid Quality Assurance Risk Assessment
Date | Risk Description | Potential Impact | Likelihood | Priority | Status | Triggering Conditions |  |  | Department | Resource requirements | Estimated Cost | Notes
2026-01-24 | Speech-to-Text accuracy degradation in noisy environments | High | High | High | Active QA | Low SNR input, overlapping speakers, accent variation |  |  | QA / AI | Noisy speech datasets, WER benchmark scripts | $120 | Target WER ≤ 15% in controlled noise tests and ≤ 20% in field tests
2026-01-24 | Hazardous-noise alert false negatives or delayed alerts | High | High | High | Mitigation Planned | Unsafe sound events >85 dB not detected within alert threshold |  |  | Embedded / QA | Calibrated dB meter, alert timing tests | $80 | Alert latency target < 2 seconds with 95% detection rate
2026-01-24 | Privacy or security breach of voice and health telemetry | High | High | High | Security Review | Weak BLE/Wi-Fi security, insecure API endpoints, exposed logs |  |  | Security | Encryption libraries, auth hardening, log review | $100 | Protect captions, biometrics, and caregiver alerts using encrypted transport/storage
2026-01-24 | Emotion detection bias or inconsistent classification | High | Medium | Medium | Model Validation | Accent, gender, speaking style, or limited training data reduce fairness |  |  | AI / QA | Labeled speech-emotion datasets and confusion-matrix analysis | $70 | Monitor class-wise precision/recall and bias across user groups
2026-01-24 | BLE / Wi-Fi connectivity loss between wearable and mobile app | High | Medium | Medium | Integration Testing | Interference, distance, reconnection failures, mobile OS background limits |  |  | Mobile / Embedded | BLE sniffer, reconnection test scripts | $60 | System should reconnect automatically and preserve queued events
2026-01-24 | Health sensor inaccuracy or calibration drift (HR / SpO2 / temperature) | High | Medium | Medium | Calibration QA | Poor sensor placement, motion artifacts, ambient conditions, firmware conversion errors |  |  | Embedded / QA | Reference sensors and calibration fixtures | $90 | Validate readings against reference devices before pilot testing
2026-01-24 | Battery drain and thermal rise during continuous AI processing | Medium | Medium | Medium | Performance Review | Always-on DSP, streaming STT, emotion inference, sensor polling |  |  | Embedded | Power profiler and thermal monitoring tools | $75 | Battery target ≥ 8 hours under representative mixed-use workload
2026-01-24 | Mobile app accessibility or readability issues for captions and alerts | Medium | Low | Low | Usability Testing | Small fonts, low contrast, poor vibration/visual cue design |  |  | UX / QA | Accessibility checklist and user testing sessions | $40 | Meet readable caption and alert UX requirements on Android devices
2026-01-24 | Firmware integration defects across DSP, sensors, and alert modules | Medium | Low | Low | Regression QA | Concurrent module updates, ISR timing conflicts, memory leaks |  |  | Embedded / QA | Hardware-in-the-loop regression suite | $85 | Run regression before every firmware release candidate
Quality Assurance Mitigation Strategies
Action for Risk Having High Priority | [5 empty] | Timeline |  |  | Resources
Priority | Task (Action Item) | Key Performance Indicators (KPIs) | Phase | Assigned | Status | Start Date | Due Date | Duration | Department | Resource requirements | Estimated Cost | Notes
High | Establish noisy-speech benchmark suite for STT and DSP tuning | WER ≤ 15% lab / ≤ 20% field; caption latency < 2s | System Testing | QA + AI Team | Planned | 2026-05-01 | 2026-05-15 | 2 weeks | QA / AI | Noise dataset pack, scripted replay harness | $120 | Run across classroom, street, cafeteria, and multi-speaker scenarios
High | Validate hazardous-noise alert path end-to-end | 95% alert detection above 85 dB; false negative rate < 5% | Integration Testing | Embedded Team | Planned | 2026-05-16 | 2026-05-30 | 2 weeks | Embedded / QA | Calibrated sound source, vibration/LED timing logs | $80 | Include missed-alert and delayed-alert boundary tests
High | Harden privacy and security controls for audio and health data | No high-severity security findings; encrypted transport/storage enabled | Security Testing | Security Team | Planned | 2026-06-01 | 2026-06-15 | 2 weeks | Security | Threat model, BLE/API penetration checklist | $100 | Verify least-privilege access, secure storage, and consent handling
Action for Risk Having Medium Priority | [5 empty] | Timeline |  |  | Resources
Priority | Task (Action Item) | Key Performance Indicators (KPIs) | Phase | Assigned | Status | Start Date | Due Date | Duration | Department | Resource requirements | Estimated Cost | Notes
Medium | Improve fairness and consistency checks for emotion detection | Class-wise F1 balanced within tolerance; no major demographic skew | Model QA | AI Team | Planned | 2026-06-01 | 2026-06-14 | 2 weeks | AI / QA | Validation datasets, confusion-matrix dashboard | $70 | Review per-emotion precision/recall and edge accents
Medium | Test app-device reconnection and offline buffering | Automatic reconnection within 10s; no critical event loss | Integration Testing | Mobile + Embedded Team | Planned | 2026-06-15 | 2026-06-30 | 2 weeks | Mobile / Embedded | BLE diagnostics, packet-loss scripts | $60 | Queue captions/alerts locally while connection is unavailable
Medium | Calibrate and verify HR, SpO2, and temperature readings | Sensor error within agreed tolerance versus reference device | Verification | Embedded Team | Planned | 2026-07-01 | 2026-07-15 | 2 weeks | Embedded / QA | Reference oximeter, temperature probe | $90 | Include motion-artifact and low-battery conditions
Action for Risk Having Low Priority | [5 empty] | Timeline |  |  | Resources
Priority | Task (Action Item) | Key Performance Indicators (KPIs) | Phase | Assigned | Status | Start Date | Due Date | Duration | Department | Resource requirements | Estimated Cost | Notes
Low | Review caption readability, alert visibility, and haptic usability | Accessibility checklist pass rate ≥ 90% | Usability Testing | UX + QA Team | Planned | 2026-05-01 | 2026-08-10 | 10 days | UX / QA | Android test devices, usability scripts | $40 | Test font size, contrast, and vibration cues with representative users
Low | Validate backup recovery for logs and historical health data | 100% retry success after reconnect in controlled outage tests | Reliability Testing | Backend Team | Planned | 2026-08-11 | 2026-08-20 | 10 days | Backend / QA | Sync retry simulator and log checks | $35 | Applies to noncritical history and analytics only
Low | Run firmware regression suite before release candidates | 0 critical regressions across DSP, sensors, alerts, and app link | Regression Testing | Embedded + QA Team | Planned | 2026-08-21 | 2026-08-31 | 10 days | Embedded / QA | HIL test bench and automated smoke suite | $85 | Gate release on regression pass
Calculations for Exposure of Each QA Risk
Range | Levels | Likelihood | Impact | Exposure | Priority
1–2 | Low | 1 | 1 | 1 | Low
3–6 | Medium | 2 | 2 | 4 | Medium
7–9 | High | 3 | 3 | 9 | High
STT Accuracy Drop |  |  | Missed Noise Alert |  |  | Privacy / Security Breach
Likelihood | 3 |  | Likelihood | 3 |  | Likelihood | 3
Impact | 3 |  | Impact | 3 |  | Impact | 3
Exposure | 3*3=9 |  | Exposure | 3*3=9 |  | Exposure | 3*3=9
Priority | High |  | Priority | High |  | Priority | High
Emotion Bias |  |  | Connectivity Loss |  |  | Health Sensor Drift
Likelihood | 2 |  | Likelihood | 3 |  | Likelihood | 2
Impact | 3 |  | Impact | 2 |  | Impact | 3
Exposure | 2*3=6 |  | Exposure | 3*2=6 |  | Exposure | 2*3=6
Priority | Medium |  | Priority | Medium |  | Priority | Medium
Battery / Thermal Issues |  |  | App Accessibility Issues |  |  | Backup Delay
Likelihood | 2 |  | Likelihood | 1 |  | Likelihood | 1
Impact | 2 |  | Impact | 2 |  | Impact | 2
Exposure | 2*2=4 |  | Exposure | 1*2=2 |  | Exposure | 1*2=2
Priority | Medium |  | Priority | Low |  | Priority | Low